## Text
Uncertainty evaluation
The mathematical model is:
Length measurements are taken during all the 3 work shifts, that is over the 24 hours, in a common working room without temperature control. The measurement procedure requires to perform 5 measurements, and to take their average as result. | Le misurazioni della lunghezza vengono effettuate durante tutti e 3 i turni di lavoro, cioè nell'arco delle 24 ore, in una sala di lavoro comune senza controllo della temperatura. La procedura di misurazione richiede l'esecuzione di 5 misurazioni, e di prendere come risultato la loro media.
Bias of the measuring instrument was evaluated by 4 measurements, showing that the average has an uncertainty of 3.7 μm with respect to the reference value.
Repeatability is declared by the manufacturer of the instrument as standard deviation of 1.3 μm.
Reading resolution of the instrument is 1 μm.
Reproducibility was tested by 10 replications made by all operators, obtaining a variance of 10.7 μm2.
Temperature sensitivity coefficient of the piece is a = 11∙10-6 °C-1 (a typical value for carbon steel, taken from handbook) and temperature sensitivity coefficient of the instrument is b = 11.5∙10‑6 °C‑1 (declared by the manufacturer), while temperature over a 24 hour time period is estimated to be 22 ± 5 °C. Measured length of the workpiece results to be 800.074 mm

## UncTable
Uncertainty table
Measuring instrument:
Factor |  | Information | Value
Bias |  | Expanded Uncertainty |  | 3.7 | μm
|  | of the average of |  | 4 | measurements
|  | based on |  | 4 | measurements
|  | at a confidence level of |  | 95%
Repeatability |  | Standard deviation |  | 1.3 | μm
Resolution |  |  |  | 1 | μm
Temperature sensitivity b |  | Declared |  | 1.15e-05 | °C-1
Measurement complex:
Factor |  | Information | Value
Reproducibility |  | Variance |  | 10.7 | μm2
Total number of measurements |  |  |  | 10
Effects contained: instrument, operators, measurand, short time environmental variations
Environmental conditions:
Factor |  | Information | Value
Temperature |  | Mean value |  | 22 | °C
|  | Variability interval | ± | 5 | °C
Effects contained: long time environmental variations
Measurand:
Length LT of the workpiece at temperature T |  |  |  | 800.074 | mm
Number of measurements averaged |  |  |  | 5
Factor |  | Information | Value
Temperature coefficient a |  | Handbook |  | 1.1e-05 | °C-1
Mathematical model
Variable xj |  |  | Statistical |  |  |  |  | Non Statistical |  | Parameters estimated
Symbol | Value | Note | Uj | Pdj | ndj | kdj | sj | aj | ka | nj | nd | nr | u²(xj) | ci=Dy/Dx | uj²(y) | uj4(y)/nj
LT | 0.80007 | Bias | 3.7e-06 | 95% | 3 | 3.1824 | 1.1626e-06 |  | 3 | 3 | 4 | 1 | 5.4068e-12 | 1 | 5.4068e-12 | 9.7446e-24
|  | Rep |  | 95% | 100 | 1.984 | 0 |  | 3 | 30 | 1 | 1 | 0 | 1 | 0 | 0
|  | Res |  | 95% | 100 | 1.984 | 0 | 5e-07 | 3 | 100 | 1 | 5 | 1.6667e-14 | 1 | 1.6667e-14 | 2.7778e-30
|  | Repr |  | 95% | 100 | 1.984 | 3.2711e-06 |  | 3 | 9 | 1 | 5 | 2.14e-12 | 1 | 2.14e-12 | 5.0885e-25
α | 1.1e-05 | Table |  | 95% | 100 | 1.984 | 0 | 2e-06 | 3 | 30 | 1 | 1 | 1.3333e-12 | -1.6001 | 3.414e-12 | 3.8851e-25
β | 1.15e-05 | Table |  | 95% | 100 | 1.984 | 0 | 2e-07 | 3 | 30 | 1 | 1 | 1.3333e-14 | 1.6001 | 3.414e-14 | 3.8851e-29
T | 22 | Estim. |  | 95% | 100 | 1.984 | 0 | 5 | 2 | 30 | 1 | 1 | 12.5 | 4.0004e-07 | 2.0004e-12 | 1.3338e-25
y=L20 | 0.80007 | m | Parameters estimated nr: Variance of y, u²(y) |  |  | 1.3012e-11 | 1.0775e-23
|  |  | Parameters estimated nr: Standard uncertainty of y, u(y) |  |  | 3.6072e-06
|  |  | Parameters estimated nr: Degrees of freedom of y, n(y) |  |  | 15
|  |  | Parameters estimated nr: Confidence level |  |  | 95%
|  |  | Parameters estimated nr: Coverage factor (Student t) |  |  | 2.1314
|  |  | Parameters estimated nr: Expanded uncertainty U(y) |  |  | 7.6886e-06 | m
Computation of Sensitivity Coefficients |  |  | Non Statistical ka: Show the main factors to be considered for reducing the uncertainty
| Dxi | xi+Dxi | Statistical Uj: y+Dy; Statistical Pdj: Dy; Statistical ndj: Dy/Dx; Non Statistical ka: Bias on LT
LT | 2.4053e-07 | 0.80007 | Statistical Uj: 0.80008; Statistical Pdj: 2.4053e-07; Statistical ndj: 1
| 0 | 0 | Statistical Pdj: -0.80007; Statistical ndj: #DIV/0!; Non Statistical ka: Show the main factors to be considered for reducing the costs
| 1.291e-08 | 1.291e-08 | Statistical Pdj: -0.80007; Statistical ndj: #DIV/0!; Non Statistical ka: Resolution on LT
| 1.4629e-07 | 1.4629e-07 | Statistical Pdj: -0.80007; Statistical ndj: #DIV/0!
α | 1.1547e-07 | 1.1115e-05 | Statistical Uj: 0.80007; Statistical Pdj: -1.8477e-07; Statistical ndj: -1.6001; Non Statistical ka: Make the hypothesis of a change in the main uncertainty factor and evaluate its effect
β | 1.1547e-08 | 1.1512e-05 | Statistical Uj: 0.80007; Statistical Pdj: 1.8477e-08; Statistical ndj: 1.6001; Non Statistical ka: Taking one half of the exp. unc. on LT, the exp. unc. on L20 is reduced to 6.0 μm
T | 0.35355 | 22.3536 | Statistical Uj: 0.80007; Statistical Pdj: 1.4143e-07; Statistical ndj: 4.0004e-07
Note on A36: Paola Pedone: simboli, valori e note delle variabili in ingresso (indipendenti)
Note on D36: Paola Pedone: tipo A scarto tipo incertezza tipo o varianza incertezza estesa di un certificato
Note on I36: Paola Pedone: si presenta nella forma di campo di variabilità risoluzione di lettura dichiarazioni di rispetto di tolleranze o conformità a norme informazioni sulla variabilità dei risultati ottenute dagli operatori
Note on D37: Expanded Uncertainty
Note on E37: Presumed confidence level
Note on F37: Presumed degrees of freedom
Note on G37: Presumed coverage factor
Note on H37: Standard deviation
Note on I37: half of variability band semiampiezza del campo di variabilità
Note on J37: B type denominator k = 3 distr. Uniforme k = 2 distr. U k = 6 distr triangolare
Note on K37: Estimated degrees of freedom in base alla attendibilità dei dati: 100 estremamente consolidata 50 molto consolidata 30 consolidata 15 da verificare
Note on L37: Estimated input multiplicity
Note on M37: Estimated output multiplicity
Note on N37: Paola Pedone: varianze delle variabili indipendenti
Note on F38: Information on the degrees of freedom (4 measurements, 3 d.o.f.) is given together with the information on expanded uncertainty.
Note on L38: The expanded uncertainty given refers to the average of 4 measurements.
Note on M38: Even if 5 measurement replications are required, the central limit theorem does not apply to systematic effects, therefore the output multiplicity should not be considered.
Note on C39: Repeatability is not used, since reproducibility is available.
Note on I42: Paola Pedone: semicampo di variabilità due unità della cifra meno significativa 2 ppm 0,2 ppm
Note on P49: Gianfranco Genta: No information on confidence level was given, therefore the conventional value of 95% is used.